## Specs and checklist
Model | Please fill
Product | Category | CHAI Specifications | Meets specifications? | Notes
Cylinders | DESCRIPTION | A container designed as a refillable cylinder used to hold compressed medical oxygen (O2) under safe conditions at high pressure (e.g. 50–200 bar). It is typically filled with oxygen when delivered from the gas supplier and includes a valve stem, an opening/closing valve, and will be graded according to size (capacity) and colour-coded to denote oxygen content. The cylinder may be made of steel, aluminium (Al) or other ferrous or non-ferrous materials and must be used together with a pressure regulator in order to release the oxygen at the correct working pressure. (SOURCE: GMDN 47225) |  | Deviations or additional descriptive elements, if applicable, please provide links to any relevant documentatio
| Technical | Refillable cylinders for compressed oxygen gas for medical use.
Specify: standard or MRI-compatible
Sizes, either ISO/US standard - specify (for example):
· "J" - 47.2 L water (6,800 L @ 137 bar)
· "G" - 23.6 L water (3,400L @ 137 bar)
· "F" - 9.43 L water (1,360L @ 137 bar)
· "E" - 4.68 L water (680L @ 137 bar)
· "D" - 2.32 L water (340 L @ 137 bar)
Nominal / working pressure: 137-150 bar
Fitted with a primary valve stem, specify standard:
· Pin-index (ISO 407/BS 850/CGA 870 valve)
· Bullnose (5/8 inch BSP (F)/BS 341 valve or CGA 540 or DIN 9 or NF 'F' or NEN Ri2)
Valve assembly material: brass
Valve handle/key operated, supplied with tools.
Material: CGA approved seamless steel/aluminum alloy/carbon fibre/other composite body
Colour coding: according to ISO/ANSI/CGA/NFPA
Supplied with valve stem cap / safety cap / valve guard
Stamped onto shoulder:
· Nominal and test pressures
· Cylinder capacity (water litres)
· Weight (kg)
· Date of manufacture
· Serial number
· Company and country of origin
· transport rating (DOT3AA or UN stamp)
· inspection agency stamp aligning with the regulatory approval
Environmental:
· Suitable for continuous operation in ambient temperature of at least 5–45 °C, relative humidity of at least 15–90% non condensing
· Capable of being stored in ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing
| Warranty | 5 years recommended
| QMS | ISO 9001 (General QMS)
| Regulatory | CE indicated by pi, "p" (Cylinders fall under TPED) or US DOT
· applies to BOTH cylinders and valves
| Product performance standards Where testing indicated, must be done by an accredited laboratory - documentation of which to be included prior to shipment | ISO 10297: Gas cylinders – Cylinder valves – Specification and type testing
ISO 14246: Gas cylinders – Cylinder valves – Manufacturing tests and examinations
ISO 7866: Gas cylinders – Refillable seamless aluminium alloy gas cylinders – Design, construction and testing
ISO 9809-1: Gas cylinders – Refillable seamless steel gas cylinders – Design, construction and testing
ISO 13769: Gas cylinders — Stamp marking
ISO 13341: Gas cylinders – Fitting of valves to gas cylinders
ISO 15996: Gas cylinders – Residual pressure valves – Specification and type testing of cylinder valves incorporating residual pressure devices
ISO 11363-1: Gas cylinders — 17E and 25E taper threads for connection of valves to gas cylinders
ISO 15001: Anaesthetic and respiratory equipment – Compatibility with oxygen
| Regional/local standards | Country-specific and regional colour gas coding and other standards apply and must be listed.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.
Extra information required: number of units.

## RFI_MFLab (hidden)
| 关于产品 | [9 empty] | 关于贵司
产品类型 Product Type | 产品标准要求 Product Specifications | 产品型号（请列出贵公司符合左侧标准要求的产品，并邮件提供电子版产品用户手册） Product Model (s) | 产品排期（假设下周下订单，预计能够供的最早出货排期至何时（日期）？ Lead time if order placed next week (date) | 本次订单所需数量 Units Required | 公司当前可出售的数量 Units Available | 法规批准/认证Regulatory Approval/Certification |  | 产品性能标准Standards, for product performance |  | 产品单价（美金） Unit price (USD) | 制造商标准 Standards, for manufacturer |  | 预付款要求 Pre-payment requirement | 公司以前有向埃塞俄比亚出口的经验吗？Does company have previous export experience to Ethiopia? | 图片示例 Picture
[6 empty] | 标准 | 如果有，请填写 “X” | 标准 | 如果有，清填写 “X” |  | 标准 | 如果有，清填写 “X”
简单氧气气面罩 Simple Face Masks | 成人式： 医用级 | LB1110, Adult(LA), 5.5mm x 2m绿色管 |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X | 生产前付清 | 有
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
不重吸面罩 Oxygen Mask with reservoir bag | 成人式： 氧气浓度：支持80%至100％的氧气浓度 管长: 2m 使用标准接头；管路（氧气管线）不弯曲 材质：聚氯乙烯（PVC）； 透明；医用级 | LB1310, Adult(LA), 5.5mm x 2m绿色管 氧袋1000ml |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
文丘里面罩（氧浓度可调面罩） Venturi Mask | 成人式：能以8至10 L / min 的流速输送低至中等流量的氧气，特定氧浓度为30％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,31%(氧气流速6L/min±10%),白色，其他符合 |  | 300 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
| 成人式：能以10至12 L / min 的流速输送低至中等流量的氧气，特定氧浓度为40％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,40%(氧气流速8L/min±10%),粉色，其他符合 |  | 300 |  | CE | X | ISO 18190
[6 empty] | FDA | X | ISO18562-1 | X
[6 empty] | 出口许可证 | X | ISO/DIS 23368
| 成人式：能以12至15 L / min 的流速输送低至中等流量的氧气，特定氧浓度为60％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,60%(氧气流速8L/min±10%),红色，其他符合 |  | 300 |  | CE | X | ISO/DIS 17256 | X
[6 empty] | FDA | X | ISO 15223-1 | X
[6 empty] | 出口许可证 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 中等大小 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带, M |  | 130 |  | CE |  | ISO 18562-1:2017 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 17510:2015 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X |  |  |  | ISO14971 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 大码。 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带，L |  | 130
气管导管 Endotracheal Tube | 管： 尺码：7。 内径：7.0mm。 外径：9.3mm。 长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 袖口： 直径：24mm | LB5010C, 带囊，ID 7.0mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X |  | ISO14971 | X
| 管： 尺码：7.5 内径：7.5mm。长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 | LB5010C, 带囊，ID 7.5mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X